## Slide 1: Ministry of higher education and scientific research Middle Technical University Baquba Technical Institute Pathological analysis Department Medical Laboratory Instrument
Assist.prof. : Eman Ismail Momammed

## Slide 2
Lecture ( 2 ):
Lecture ( 2 ):
Condenser of microscope
It is a lens that used to concentrate light from the illumination source that is in turn focused through the object and magnified by the objective lens .
There are three types of condenser:
1-Abb condenser:.It has two controls :
- Moves the Abb condenser closer to or further from the stage .
-The iris diaphragm ,which controls the diameter of the beam of light .

## Slide 3
A diaphragm: is a thin opaque structure with an opening ( aperture ) at its center .
Abb condensers are mainly vital for magnifications of above 400X . It is considered as the simplest type

## Slide 4
2-Aplanatic condenser: It corrects for spherical aberration in the concentrated light path , spherical aberration is an optical effect observed in an optical device ( lens , mirror ) that occurs due to the increased refraction of light rays when they strike a lens or a reflection of light rays when they strike a mirror near its edge in comparison with those that strike nearer the center .
Plan lenses can reduce or eliminate spherical aberration.

## Slide 5
Spherical aberration

## Slide 6
3-Achromatic condenser: There are corrected for chromatic aberrations. Achromatic lens is a lens that is designed to limit the effects of chromatic and spherical aberration. Chromatic aberration is a type of distortion in which there is a failure of a lens to focus all colors to the same convergence point .
Chromatic aberration occurs due to the variation of refractive index with wavelength for a lens material

## Slide 7
This wavelength dependence results in slightly different focal lengths for different wavelengths of light. Compound lenses, called achromats, can reduce or eliminate chromatic aberration because the components are chosen such that the variation in refractive index as a function of wavelength cancels out.
Chromatic aberration

## Slide 8
Numerical Aperture (N.A.):
This is a number that expresses the ability of a lens to resolve fine detail in an object being observed. It is derived by a mathematical formula (n sine θ ) and is related to the angular aperture of the lens and the index of refraction of the medium found between the lens and the specimen. The physical size of the lens is important in determining the N.A. of the lens
NA = N sin θ

## Slide 9
Where:
n : is the index of refraction of the medium in which
the lens is working .
(1.00 for air, 1.33 for pure water and 1.52 for immersion oil )
θ : is the maximal half-angle of the cone of light that can enter or exit the lens.
Resolution power (Resolving) : It is a capacity to distinguish two adjacent point depend on :
1- numerical aperture of lens .
2- wave length of used light .

## Slide 10
Note :
- The human eye can separate dots ( 0.25 mm apart.)
-The light microscope can separate dots (0.25 micron apart.).
- The electron microscope can separate dots less than angstrom .

## Slide 11
Light path in the Microscope