## Slide 1: Ministry of higher education and scientific research Middle Technical University Baquba Technical Institute Pathological analysis Department Medical Laboratory Instrument
Assist.prof. : Eman Ismail Momammed

## Slide 2
Lecture ( 5 ) :
Lecture ( 5 ):
Photometry
Visible light spectrum:
When a beam of light passes through a colored solution, it interacts with matters in the solution and the result may be refraction, reflection, absorption and transmission among others
-Refraction: Is defined as sudden change in the direction of the beam when the light passes from one medium to other with a different physical density.

## Slide 3
-Reflection: Is a condition where the beam returns
back towards its source such as mirror.
-Absorption :Is a situation where some components of the light (colors) are retained or absorbed.
-Transmission: Refers to the situations where some portions of the light permitted to pass through a given medium .

## Slide 4
Radiation is characterized by waves on which basis the electromagnetic radiation spectrum could be divided in many regions including gamma rays, x-rays, ultraviolet rays, visible light , infrared, microwaves and radio waves . visible region is the radiant energy to which the human eye responds and their wavelength varies between 400 and 700 nm

## Slide 5
Wavelength of about 700 nm are seen by the eyes as red colors while those of progressively shorter wavelengths give in descending order to orange, yellow, green, blue, indigo and finally violet colors which is produced in the short wavelength of 400nm
Visible light spectrum

## Slide 6
electromagnetic spectrum

## Slide 7
Beer’s and Lambert’s Law:
Most colorimetric analytical tests are based on the Beer’s - aLambert’s law which states that under the correct conditions the absorbance of a solution when measured at the appropriate wavelength is directly proportional to its concentration and the length of the light path through the solution. Using a standard, this law can be applied to measuring the concentration of a substance in unknown (test) solution by using the formula:

## Slide 8
Concentration of test ( Ct ) =

## Slide 9
Photometer
It is instrument measure the absorbance optical density . This equipment will measure the transmission and optical density depending on color filter which give the complement color and in this case the reading less accurate than spectrophotometer which is use monochrometer

## Slide 10
Parts of photometer :
1-Light source .
2- Filter : to give approximate wave length according to the color .
3-Sample holder: which can hold the sample .
4- Photocell : it will convert the light to electrical current
5- Galvanometer : which can measure the current from photocell.
6-Zero adjustment : which can adjust the zero point and reading .

## Slide 11
Types of filter :
1- Blue filter : it will pass the wavelength between ( 400 – 495 ) nm .
2- Green filter : it will pass the wavelength between( 500 – 580 ) nm .
3-Red filter : it will pass the wavelength between ( 600 – 800 ) nm .
Photometer device

## Slide 12
Light path in the photometer