## Slide 1: Ministry of higher education and scientific research Middle Technical University Baquba Technical Institute Pathological analysis Department Medical Laboratory Instrument
Assist.prof. : Eman Ismail Momammed

## Slide 2
Lecture ( 8) :
Centrifuges
It is a device for separating two or more substances from each other by using centrifugal force. Centrifugal force is the tendency of an object traveling around a central point to continue in a linear motion and fly away from that central point.
Centrifugation can be used to separate substances from each other because materials with different masses experience different centrifugal forces when traveling at the same velocity and at the same distance from the common center.

## Slide 3
Principles work of centrifuges:
The basic physics work of centrifuges is gravity and generation of the centrifugal force to sediment different fraction . Centrifugal field ( G ) depend on :
1-Angular velocity ( w ) in radians/ sec .
2-Radial distance ( r in cm ) of particle from axis of rotation .
G = w2 r

## Slide 4
Rate of sedimentation depends on :
Mass of particle ( density and volume )
Density of medium .
Shape of particle .
Friction .

## Slide 5
Types of centrifuges:
There are many types of centrifuges, but the basic principle is the same, all use centrifugal force:
1- Desk top centrifuges ( Low speed centrifuges ) :
- It is very simple and small.
- maximum speed of 3000 rpm.
- Don't has any temperature regulation system .
- Used normally to collect rapidly sedimenting substance such as blood cell

## Slide 6
Head of centrifuge

## Slide 7
2- High speed centrifuges:
-Maximum speed of 25000 rpm.
-Temperature maintained at 0 – 40 C by thermocouple. - Used to collect microorganism, cell, large cellular organelles.
-Useful in isolating sub cellular organelles ( nuclei , mitochondria and lysosomes )

## Slide 8
3- Ultracentrifuges :
- Operate at speed of 75000 rpm .
- Rotor chamber is sealed and evacuated by pump to attain vacuum .
- Refrigeration system temperature 0 – 40 C .
There are two types of ultracentrifuge :
A- Analytical centrifuge.
B-Preparative centrifuge.

## Slide 9
Parts of centrifuges:
1-Speedmeter : increasing and decreasing of speed .
2- Timer : control of time .
3- Speed indicator .
4- Head with sample holder .
5- Cover .
6- Electric motor( attached head for carrying sample ).
Power for centrifuges:220 – 240V

## Slide 10
Applications of centrifuges :
1- Remove cellular elements from blood to provide cell free plasma or serum for analysis .
2- Isolate chemical precipitated protein from an analytical specimen .
3- Separate protein bound from free ligand in immunochemical .
4-Separate lipid components .

## Slide 11
Use and care of centrifuges:
1. Reading the manufacturer’s instructions.
2. Placing a centrifuge on a firm level bench out of direct sunlight, towards the back of the bench.
3. Whenever possible using plastic tubes made from polystyrene or autoclavable.
4. Closing the centrifuge top before turning it on.
5. Always balancing the tubes that are being centrifuged, tubes of the same weight should be placed directly opposite to each other. Tubes should also be of the same size and should also contain the same amount of liquid..

## Slide 12
6. Increasing spinning speed gradually is important. That is if you are required to spine a mixture at 3, 000 rpm, first put the dial on 1,000 rpm.Give the centrifuge a chance to come up to that speed, then turn up the dial a little further until it reaches the desired 3,000 rpm. Five minutes are the usual time required to centrifuge most substances.
7. Never open the centrifuge while it is still spinning.